National Institutes of Mental Health (NIH)
 9000 Rockville Pike
 Bethesda, Maryland 20892

Phone: 301/402-2362

MEG CORE fACILITY

MEG

Data  Acquisition Tutorial

Version 6/03

MEG Core Facility

# MEG Data Acquisition Tutorial

National Institutes of Mental Health (NIH)
 9000 Rockville Pike
 Bethesda, Maryland 20892

Phone: 301/402-2362

http://kurage.nimh.nih.gov/meglab

*This manual contains information and images from CTF Systems Inc. MEG/EEG Operation and Technical Reference Manual & Omega 2000 Software User Manual, Version 4.14

| # Table of Contents |  |  |  |  |
| --- | --- | --- | --- | --- |
| # Instructions |  |  |  |  |
| # Section I SUBJECT PREP | # Page |  | # Section V I # SUBJECT MONITORING # INSIDE THE MSR | # Page |
| - Demetaling | 5 |  |  |  |
| - Degaussing | 5 |  | # Video Display Monitor | # 39 |
|  |  |  | # B. Voice Intercom / Parabolic Speaker | # 40 |
| Section I I MEG SETUP |  |  | - Infrared Lighting | 40 |
|  |  |  | # D. Video Cameras | # 40 |
|  |  |  | # E. VCR | 41 |
| - Head Size / Fit Test | 6 |  |  |  |
| - Fiducial Points / Head Coil Placement | 6 |  | # Section V I I RECORDING DATA |  |
| - Electroculogram Eye Lead Placement | 7 |  |  |  |
| Section I I I EEG SETUP |  |  | A. Electronics / System Characteristics | 43 |
|  |  |  | - ADC Setup | 44 |
|  |  |  | B. Data Acquisition / Collection Parameters | 45 |
|  |  |  | C. Acquisition Dialog | 45 |
| - Head Size / Fit Test | 8 |  | D. Channel Selection | # 46 |
| - Fudicial Points / Head Coil Placement | 8 |  | E. MEG System EEG Impedance Check | # 46 |
| - Electroculogram Eye Lead Placement | 9 |  | F. Head Localization | # 46 |
| - Reference Electrodes Placement | 10 |  | # G. Pushing Data to the MEG Server | # 48 |
| - EEG Easy Cap System | 10 |  | - Transferring Data to DVD | 50 |
| 1. Marking Frontal, Cz & Occipital Locations | 10 |  | I. Transferring Data to a Remote | 51 |
| 2. Mounting the Cap | 12 |  | Workstation |  |
| - Digitizing Electrode Positions | 13 |  |  |  |
| - Reducing Skin Impedance | 15 |  |  |  |
| H. Applying Electrolyte | 15 |  |  |  |
| I. Electrode Impedance Meter Check | 16 |  | Section V I I I CLEAN UP PROCEDURE |  |
| J. Electrode Jack Box | 17 |  |  |  |
| Section I V # STIMULUS / RESPONSE EQUIPMENT |  |  |  |  |
|  |  |  | - MEG |  |
|  |  |  | 1. Subject Clean Up | 53 |
|  |  |  | 2. Equipment Clean Up | 53 |
|  |  |  | - EEG |  |
| # A. Visual Stimulus | # 18 |  | 1. Subject Clean Up | 54 |
| # LCD Projector Setup | # 19 |  | 2. Disinfection Procedure | 54 |
| 2. Optical Sensor System | 23 |  | 3. Equipment Clean Up | 54 |
| 2. Video Stimulus Presentation / Eye | 24 |  |  |  |
| # Tracking System |  |  |  |  |
| # B. Auditory Stimulus | 25 |  | Section V I I I APPENDIX |  |
| # C. Somatosensory / EMG Stimulus | 26 |  |  |  |
| # D. Response Pads | 28 |  |  |  |
| 1. Lumitouch | 28 |  |  |  |
| # 2. FORP | 28 |  | Variance Reporting | 55 |
| # 3. NeuroScan | 28 |  |  |  |
|  |  |  | INDEX |  |
| # Section V # MSR & SUBJECT SUPPORT |  |  |  |  |
| - Room Lighting | 29 |  |  |  |
| - Door Operation | # 29 |  |  |  |
| C. Positioning a Subject in the Chair | 31 |  |  |  |
| - Head Coil Cable Module | 31 |  |  |  |
| 2. Mounting the Jack Box to the | 31 |  |  |  |
| Amplifier |  |  |  |  |
| 3. Stabilizing Head cuff | 32 |  |  |  |
| D. Chair Adjustments | 33 |  |  |  |
| E. Installing the Bed for Supine Recording | 36 |  |  |  |
| F. Positioning a Subject on the Bed | 38 |  |  |  |

Instructions

This manual is designed to provide the user with didactic instructions for basic MEG / EEG setup and acquisition.  It is intended to be a comprehensive but not exhaustive.

It is organized into major sections and then further divided into subsections in the order that a study might logically be conducted.  Most sections are designed to provide step-by-step instructions to guide the user through the setup or acquisition process while other sections are intended to provide some general information about a procedure, process or equipment.  In addition, there are notes interspersed throughout the text to provide the user with helpful hints, useful tips, warnings and /or precautions.  Links are also included to provide more detailed information.

It is recommended that the new user start at Section I and proceed systematically through each section to the end.  Some steps are optional depending upon the type of study that is needed or required.  The more experienced user may wish to use this manual as a reference and may choose to skip from section to section as desired.

Subject Prep

De-Metaling

Before entering the MSR subjects must remove any metal from their person.  There are no known risks associated with MEG recordings.  Although wearing metal objects while in the MSR is not harmful to the individual, bringing metallic  or magnetic objects into the MSR  causes interference and high noise levels in the recording.

Subjects must remove:

- Bras with metal under wires

- Makeup - mascara, eye liner, eye shadow

- Clothing containing metallic shiny threads or glitter

- Metal on clothing i.e., metal buttons, snaps or trimming

- Belts

- Keys

- Watches

- Rings or other jewelry

- Coins

- Eyeglasses

- Cell phones

- Pagers

- Credit Cards

Non-Magnetic Clothing

Small metal objects such as hair pins, paper clips, safety pins, etc.

Additionally, everyone entering the MSR must also remove their shoes to prevent magnetic

dust from getting into the MSR.

It is recommended that subject are screened and given instructions before arriving for testing.  However, if subjects arrive with metal on their clothing that can not be easily removed they may be issued non-magnetic (paper) clothing which will be provided by staff.

Degaussing

If a subject arrives with embedded metal on their person, i.e., a permanent retainer, which can not be removed it may be necessary to de-magnetize the subject.  Care should be taken not to use the Degausser close to the computer monitors as it may permanently damage the screens. The Degausser  must NOT be used in the MSR.

- Plug in the Degausser.

- Move 3-4 feet away from all computer monitors.

- Pass the Degausser in a circular motion in front of the area of interest 3 or 4 times.

Degausser

- Unplug Degausser.

*It is recommended that the subject not have had an MRI within 3-4 days before having a MEG scan because of the possibility of residual magnetization.

MEG Setup

# Head Size/Fit Test

# Used to determine whether the subject will fit the MEG sensor helmet. The helmet will fit subjects within 98 percentile for head size.

- Determine head size by measuring the:

  - Head Circumference,

  - Nasion-Inion Distance and

  - Preauricular Point to Preauricular Point Distance.

- Determine fit by placing the MEG helmet along with a protective cap over the head.

Helmet Used for Fit Test

Record measurements.

| Gantry Limits | Nasion-Inion | Preauricular | Circumference |
| --- | --- | --- | --- |
| MEG / EEG | ≤ 40cm | ≤ 40cm | ≤ 60cm |

# Fiducial Points/Head Coil Placements

# Used for head localization (the process of locating the subject’s head in space relative to the Dewar). Used later to co-register the MRI Coordinate System with the MEG Head Coordinate System.

If the subject has had previous digital pictures of the fiducial points retrieve images to match placement.

Left Preauricular Point

Nasion

Right Preauricular Point

- Place a small black dot at the nasion (indentation between the forehead and the nose) midline. *The dot may be placed slightly higher if the nasion curvature is such that it does not allow level placement of the head coil.

- Place a small black dot below the left & right preauricular points centered (about 3-4cm) in front of the tragus. This may be measured for accuracy.

- Take a picture of these locations using the HP 715 Digital Camera.  Download and print the pictures.

- Attach Small (MEG) Adhesive Washers to head localization coils.

- Attach head localization coils at the fiducial marks.  *Make sure the marks line up with the center of the head coils.

| Head Coil Cables |  |  |
| --- | --- | --- |
| Nasion | Right Ear | Left Ear |
| Blue | Red | Green |

- Place tape over the cables to secure.

Head Coil  Placement

Electroculogram (EOG) Eye Lead Placement

Used to identify and monitor eye blinks and saccadic eye movements.

Left Infra Orbital (LIO) - Place electrode on the orbital ridge centered directly under the eye.

Left Outer Canthus (LOC) - Place electrode at the lateral junction of the upper and lower eye lid.

- Clean the skin on the cheek near the eyes with alcohol and Nuprep Skin Cleaner.  Wipe the skin to remove the gritty residue.

- Attach the Large EEG Adhesive Washers to the electrodes.

- Place the Infra Orbital electrode.

- Place the Outer Canthus electrode.

EOG Placement

Press the electrodes onto skin.

- Apply Elifix Electrolyte through the electrode opening.

- Check the impedances using the Grass Impedance Meter.

  - Plug in the left infra orbital electrode in jack 1 and the left outer Canthus electrode in jack 2.

  - Push the red ON button.

  - The Electrode Selector should be at 1 or 2.

  - EOG impedance should be < 50k ohms.

8.    Secure with tape.

*Refer to section entitled “Electrode Impedance Meter Check” for more detailed information

on reducing impedances.

EEG Setup

# Head Size/Fit Test

# Used to determine whether the subject will fit the MEG sensor helmet. The helmet will fit subjects within 98 percentile for head size. *This number is decreased slightly when using the EEG Cap.

- Determine head size by measuring the:

  - Head Circumference,

  - Nasion-Inion Distance and

  - Preauricular Point to Preauricular Point Distance.

- Determine fit by placing the MEG helmet along with a protective cap over the head.

Helmet Used for Fit Test

Record measurements.

| Gantry Limits | Nasion-Inion | Preauricular | Circumference |
| --- | --- | --- | --- |
| MEG / EEG | ≤ 40cm | ≤ 40cm | ≤ 60cm |

# Fiducial Points/Head Coil Placements

# Used for head localization (the process of locating the subject’s head in space relative to the Dewar).  Used later to co-register the MRI Coordinate System with the MEG Head Coordinate System.

If the subject has had previous digital pictures of the fiducial points retrieve images to match placement.

Left Preauricular Point

Nasion

Right Preauricular Point

- Place a small black dot at the nasion (indentation between the forehead and the nose) midline. *The dot may be placed slightly higher if the nasion curvature is such that it does not allow level placement of the head coil.

- Place a small black dot below the left & right preauricular points centered (about 3-4cm) in front of the tragus. This may be measured for accuracy.

- Take a picture of these locations using the HP 715 Digital Camera.  Download and print pictures.

- Attach Small (MEG) Adhesive Washers to head localization coils.

- Attach head localization coils at the fiducial marks.  *Make sure the marks line up with the center of the head coils.

| Head Coil Cables |  |  |
| --- | --- | --- |
| Nasion | Right Ear | Left Ear |
| Blue | Red | Green |

- Place tape over the cables to secure.

Head Coil Placement

Electroculogram (EOG) Eye Lead Placement

Used to identify and monitor eye blinks and saccadic eye movements.

Infra Orbital (LIO) - Place electrode on the orbital ridge centered directly under the left eye.

Outer Canthus (LOC) - Place electrode at the lateral junction of the upper and lower left eye lid.

- Clean the skin on the cheek near the eyes with alcohol and Nuprep Skin Cleaner.  Wipe the skin to remove the gritty residue.

- Attach the Large (EEG) Adhesive Washers to the electrodes.

- Place the Infra Orbital electrode.

EOG Placement

Place the Outer Canthus electrode.

- Press the electrodes onto skin.

- Apply Elifix Electrolyte through the electrode opening.

- Check the impedances using the Grass Impedance Meter.

- Plug in the left infra orbital electrode in jack 1 and the left outer Canthus electrode in jack 2.

- Push the red ON button.

- The Electrode Selector should be at 1 or 2.

- EOG impedance should be < 50k ohms.

8.    Secure with tape.

*Refer to section entitled “Electrode Impedance Meter Check” for more detailed information on reducing impedances.

Reference Electrodes / Mastoid Placement

Used to define the electronics common point.

Left and Right Mastoid Placement - Bone prominence behind the ear.

- Clean the skin behind the left and right ear with alcohol and Nuprep Skin Cleaner.  Wipe the skin to remove the gritty residue.

- Attach Large (EEG) Adhesive Washers to the electrodes.

- Press the electrodes onto the skin.

- Apply Elifix Electrolyte through the electrode opening.

Mastoid Placement

- Check the impedances using the Grass Impedance Meter.

a.    Plug in the left mastoid electrode in jack 1 and the right mastoid electrode in electrode jack 2.

b.    Push the red ON button.

c.    The Electrode Selector should be at 1 or 2.

d.   Mastoid electrode Impedance should be less than or     equal to 5K ohms.

*Refer to the section entitled “Electrode Impedance Meter Check” for more detailed information on reducing impedances.

- Secure with tape.

EEG Easy Cap System

The Easy Cap is a modular EEG-recording cap system with 64 ring adaptor locations using Ag/AgCl sintered electrodes.  *Other application methods may be used, however, careful attention should be given to the choice of electrodes; refer to Electrode Criteria.

Measuring / Marking Frontal Polar, Cz & Occipital Locations

- Measure the nasion – inion distance. Place a mark at 50% of the total distant. This marks Cz.

- From the nasion place a horizontal mark at 10% of the nasion – inion distance.  This marks FPz.

- Place a vertical mark on the forehead midline nasion.  For accuracy measure the distance between the eyebrows; mark midway crossing the horizontal mark. This completes Fpz.

- From the inion place a horizontal mark at 10% of the nasion - inion distance.  This marks Oz.

- Measure head circumference from the Fpz position. Place a vertical mark at 50% of the distance of the circumference.

This completes Oz.

- Measure and mark 5% of the circumference distance on the left of FPz (on the same horizontal plane).  This completes Fp1.

- Measure and mark 5% of the circumference distance on the right of FPz (on the same horizontal plane). This completes Fp2.

- Measure and mark 5% of the circumference distance on the left of Oz (on the same horizontal plane).  This completes O1.

- Measure and mark 5% of the circumference distance on the right of Oz (on the same horizontal plane). This completes O2.

- Measure from left to right preauricular points. Mark the midpoint.

This completes Cz.

Fp1, Fpz, Fp2 Locations 
Mounting The Cap

- Select a cap size based upon the head circumference measurement:

| Small | Med | Large |
| --- | --- | --- |
| ≤ 56cm | ≤ 58cm | ≤ 60cm |

- Rotate the adaptors in the cap so that the narrow opening is toward the back of the head.

- Place the Large (EEG) Anchoring Adhesive Washer on the inside of the cap on the adaptors at Fp1 & Fp2   locations.

- Press Fp1 & Fp2 with the attached adhesive washers onto the forehead at the marked locations.

- Have the subject hold Fp1 & Fp2 in place while turning the cap over the head.

- Gently pull the cap over the head and ears into place.

- Check and adjust the cap so that the Cz electrode is at the Cz mark.

- Check and adjust Fp1 & Fp2 and O1 & O2 to assure that they are placed correctly; if not, choose another cap size.

EEG Cap Placement

Anchor the cap with the chest belt (or the chin strap).

- Snap the electrodes into the adaptors on the cap in a systematic manner starting at the back of the head.  *This step isn’t necessary if the electrodes are already on the cap.

- Insert the electrodes so that the lead wire points toward the narrow side of the adaptor.

- Using an old ballpoint pen (or the fingers) where the mine (the point of the pen) has been removed push the electrode into the adaptor.

Do not to bend or place excessive pressure on the lead wire where it attaches to the electrode.

Digitizing Electrode Positions

The Polhemus Digitizer is used for the digitization of the EEG electrode positions and for the determination of head shape. Digitize all electrode positions using the Polhemus.

| Polhemus Receiver Module Polhemus Transmitter Module Stylus Pen | - After the Easy Cap has been placed on the subject, position the small Receiver Module on the back of the head on top of the cap near midline. Secure with tape. - Place the large Transmitter on a non-metallic stand midline directly behind the subject. Tape the transmitter to the stand to secure. - Turn on the laptop computer located on the portable cart. - Open the Polhemus Electrode Digitizing Program: - Double click on the Electrode Digitizer Program. - The software should detect two (2) stations: the receiver and transmitter. - If the software does not recognize both stations, check the connections then restart the program. - Define Head Coordinates: - Click on Define Head Coordinates. - Define the head coordinates by placing the stylus pen tip at the center of a head coil location (nasion, left preauricular, right preauricular) then press the stylus button. If a mistake is made, - Press the stylus pen until all head coordinates have been entered then, - Press define Head Coordinates again. Do not save the previously acquired locations simply start over from the beginning. |
| --- | --- |

|  |  | - Digitize Electrode Locations: - Click on Digitizing Start / Stop. - Place the stylus pen at the center of an electrode location as designated by the digitizing program then press the stylus button. - A single data point is collected each time the stylus button is depressed and a beep will alert the user to each acquired point. - The program will assign a number to the electrode position in the sequence the electrodes are selected, i.e., 1, 2, 3, 4, etc. - If a mistake is made, click on Remove Last Point and try that position again. - Click on Digitizing Start /Stop when all electrode locations have been done. - At the end the program a head diagram will appear on the screen. - If the diagram is: - Correct then continue with Step 8. - Not correct then check to make sure that the receiver or transmitter did not move then repeat digitization (steps 6a –d). - Click on File / Save Data. - Save the data to the hard drive and - Save to a floppy disk. *The filename is the current date. - Exit the program. Eject disk. Turn off laptop. - Transfer the Polhemus head coordinate file from the disk to the computer at the acquisition workstation / console. |
| --- | --- | --- |
| If performing EEG proceed to next section “Reducing Skin Impedance & Applying Electrolyte” otherwise skip to the appropriate Stimulus Equipment section. |  |  |
|  | The operator and anyone within 10 feet of the transmitter and receiver should de-metal. The general rule of thumb is three times the separation distance between the transmitter and receiver from any metal in the room. For example, if the separation distance is one foot, metal within three feet of the transmitter or receiver may cause distortion and inaccurate measurements. |  |

Reducing Skin Impedances

Nuprep is intended for to reduce skin impedance. It is a topical abrasive used to clean the skin to remove any natural or applied oil, lotions and/or dead skin.  Remove with water.

  - Insert the wooden end of a cotton swab through the opening of the electrode to push hair aside (in one direction) until the skin is clearly visible.

  - If the hair is long, lift up the electrode a little and let it down again while moving back the swab to grab a new bunch of hair (this will prevent the hair from slipping back).

  - Dip the swab into the Nuprep (Abralyt or other skin prep) and insert through the electrode opening.

  - Twirl the swab against the skin using the thumb and index finger.

Nuprep Skin Prep

Applying Electrolyte

Elifix Conductive Cream is a white opaque, water soluble electrolyte used as a conductive medium.  Remove with warm water.

      - Fill a syringe with Elifix (Abralyt or other electrolyte).

      - Squirt Elifix into the middle of each electrode opening, completely filling the electrode space (the space from the skin to the electrode).

Elifix Electrolyte

When ELECTROLYTE: If the electrodes are closely spaced, draw back the syringe concurrently to prevent electrolyte from swelling underneath the adaptor rim and bridging to the adjacent electrodes.

Electrode Impedance Meter Check

The Grass S88 Impedance Meter measures AC impedance (resistance and capacitance).  An impedance check should be performed on all electrodes used for EEG or EMG.

When two or more electrodes are plugged into the electrode switch positions, all electrodes except the one selected electrode being measured are connected together and provide the reference for the one being measured.

- The Electrode Selector Switch should be at position 1 or 2.

- Plug the ground electrode (FPz) into electrode jack 1 position.

- Plug the other electrodes into the jack 2 position in a systematic manner one at a time.

- Press the Red ON Button to activate the impedance check.  The power will automatically shutoff after approximately 2 minutes. (There is no OFF button.)

- If the impedance is too high -  take a cotton tip swab dipped in Nuprep (Abralyt or other skin prep) and twirl against the skin until impedance is at an appropriate level. Then re-insert Elifix electrolyte.

| Impedance Recommendations |  |
| --- | --- |
| EEG Electrodes | 5Kohms |
| EMG Electrodes (EOG, ECG, etc.) | 50Kohms |

- After checking the impedance of each electrode plug the electrode into the appropriate electrode jack box position.

Impedance Meter

If the electrode impedance exceeds 199.9K ohm, a “1” will appear in the left display window with all other digits blanked out indicating the impedance is too high. Take a cotton swab that has been dipped in Nuprep and insert it through the electrode opening and twirl against the skin. Re-apply electrolyte. Re-check the impedance.

## Electrode Jack Box

Ground Electrode (Z) – Reference Electrode (Ref) Static Control Button Bipolar Channels Unipolar Channels Electrode Jack Box Mounted On The Amplifier

| ELECTRODE | JACK LOCATION |
| --- | --- |
| FPz | Ground (green); Z position. |
| Linked Mastoids | Ref position (red). |
| EOG – Left Outer Canthus EOG – Left Infra Orbital | A+ (black); Bipolar Channel A- (blue); Bipolar Channel (EEG Channel 057) |
| Cerebral Electrodes | Place according to selected montage; Numbered jacks; Unipolar Channel. |

If Using EOG:

      - Plug into a bipolar channel.

If Performing EEG:

        - Plug FPz into the ground (green / Z) electrode position.

Linked Electrode

Adaptor

Plug the Mastoids into a Linked Electrode Adaptor (Y Connector) then place the single end of the adaptor into the Ref (red) position to link the mastoids together.

        - Plug the cerebral electrodes into the appropriate MEG System Jack Box unipolar connection according to the particular montage design.

*Care should be taken to make sure to plug electrodes in the correct jack box location.  Check the montage.  Electrode labels on jack box are in the lower right corner of each jack position.

Proceed to the Magnetically Shielded Room (MSR).

# Stimulus / Response Equipment

*The stimulus should be set up prior to the subject’s arrival.

Visual Stimulus

The visual software includes:

- Stim – a DOS based software program which delivers visual stimuli.

- Presentation – a stimulus delivery and experimental control software system for neuroscience. Runs on Windows and uses standard PC hardware. Allows more precise timing for stimulus delivery and experimental control. Provides millisecond temporal precision and complete timing information for all stimulus and response events. Features include: - Images stored as bmp, jpg, or pcs (picture stimuli). - Use any adapter display mode from 32 bit color to palletized display modes. - Synchronizes stimulus appearance with monitor vertical scan (Picture Timing Control). - Shows a new stimulus up to every vertical scan. - AVI and MPEG video (video stimuli). - Animations. - Built-in programming language for custom control (PCL). - Controls stimulus presentation in real-time. - Adjusts stimuli based on responses or external inputs in real-time. - Reports times for any event of interest (logfiles). - Verifies all timing to detect operating system problems.

The visual interface includes the:

- Sharp LCD Projector System

- Avotec Silent Vision Binocular Fiber Optic Video Stimulus Presentation System

- Avotec Real Eye Tracking System

LCD Projector System

The LCD Back Projection System consists of a projector and a screen mounted to the gantry. Images are projected through a hole in the shielded room wall into a set of mirrors which projects the image onto the screen.  The image size is up to 40 degrees visual field for eye-screen distance of 37.2 to 52.5 cm.

The Sharp LCD Projector System includes:

- LCD Projector

- Screen

- Mirrors

- Remote Control

Sharp LCD Projector

- Turn on the power to the projector by pressing the POWER ON touchpad which is located on the under side of the projector.  The power indicator will turn green.

- If the power is turned off and then immediately turned on again, it may take a short while before the lamp turns on.

- STANDBY MODE:  If OFF is pressed twice, the POWER indicator will light up red and the cooling fan will run for about 90 seconds.  The projector will then enter standby mode.  The power can be turned on again by pressing ON.  When the power is turned on, the POWER indicator and the LAMP REPLACEMENT indicator will be green.

- WARM UP MODE: The flashing green LAMP INDICATOR shows that the lamp is warming up.  Wait until the indicator stops flashing before operating the projector.

- To turn the power OFF press the off button twice, it will enter the standby mode (the power indicator will light red) and after a short delay the power will turn off.

**The LCD Projector has a 30-40 ms delay before it displays an image.

In order to preserve the bulb life, turn the projector off after each use.

Positioning the Projector:

The projector position may be adjusted manually by shifting slightly to the left or right and up or down.

Lens Adjustments:

Manual Adjustment

- Images may be focused by turning the silver and black lens collar located on the projector.

- The projector lens should be centered in the middle of the screen and should be perpendicular to the mirror; otherwise the image will be distorted, making viewing difficult.

Program Adjustment

This function can be used to adjust the focus, zoom, and keystone settings. *Generally only manual adjustments are necessary.

- Press LENS to select mode.  Each time LENS is pressed, the screen changes.

- Press ENTER to display test pattern.

- Press the arrow buttons to make adjustments.

- Press LENS until normal screen appears.

- To reset the KEYSTONE setting, press UNDO.

Mirror Adjustments

Wall Mirror Adjustments:

- Project an image onto the mirror.

- Adjust the wall mirrors:

a.  Position the top mirror in front of the lens so that the image is reflected down to the bottom mirror.

b.  Adjust the bottom mirror so that the image is

projected centered onto the screen.

Top Mirror

Ceiling Mirror Adjustments ( for Supine Recording):

- Place the ancillary wall mirror on the 2nd set of wall brackets.

- Project an image onto the mirror.

- Slide the mirror back or forward so that the image is projected down onto the screen.

Bottom Mirror

Screen Adjustments:

- Position the screen so that the image is projected centered onto the screen. If possible the screen should be perpendicular to the floor.

- The screen distance can be increased or decreased by loosening the black knob located on the telescoping screen arm and adjusting the arm length.

Screen inside the MSR

If necessary, the screen can also be removed and flipped so that it is closer to the subject. Loosen the knob located on the telescoping screen arm.  Pull the arm out, flip the screen and re-insert.  Position the screen and tighten the knob.

- Use the black knob located on the Gantry Mounting Bar to raise and lower the screen.

Black Screen Function:

This function can be used to superimpose a black screen over the projected image. The Remote Control or the Touch Pad located under the projector can be used to activate this function.

- Blacking out the Projected Image WITH On-Screen Display:

- Press BLACK SCREEN.  The screen turns black and “BLACK SCREEN” is displayed on the screen.

- To return to the original projected image, press BLACK SCREEN again.

For remote control operation, slide the MOUSE/ADJUSTMENT switch to the MOUSE position.

LCD Projector

Remote Control

Optical Sensor System

The optical sensor is positioned in front of the projector.  It measures the light output of the LCD projector and is used to determine accurate timing of visual stimulus onset times.

The system includes:

- Optical Sensor

- LCD Projector

- Illumination Sensor Box

Illumination Sensor Box

- Place the Power Switch for the Illumination Sensor Box which is located on the stand outside of the MSR in the ON position.

- Adjust the Optical Sensor located in the projector tube. *Sensor may be taped to the lower mirror for better accuracy.

- Check the rack to ensure that the optical sensor output is plugged in. It is typically on ADC16.

Video Stimulus Presentation / Eye Tracking System

The Avotec Silent Vision Video Stimulus Presentation / Real Eye Tracking System allows simultaneous visual stimulation and eye tracking for MEG applications. The Avotec Binocular Fiber Optic Video Presentation System consists of an optic device for attachment to the MEG system gantry.  The Real Eye Tracking System is an infrared eye position tracking system which takes a picture of the eyes.

It consists of four subsystems:

- Avotec Silent Vision - Video presentation system

- Eye Illumination Sensor – infrared eye illumination

- Avotec Real Eye Tracker Interface – video based eye tracking system

- Eye Tracking Computer Interface (currently not available)

Silent Vision Infrared System

  - Press the Power button for the Silent Vision which is located at the operator’s console.

  - Press the Power Switch for the Real Eye Tracker Interface located at the operator’s console.

  - Turn the Real Eye – Infrared Level Knob (IR) clockwise until (from high to low) approximately one-third of the LED indicator is lit.

  - The monitor of the Real Eye Tracker Interface should show an eye image.

  - If the eye image is rotated away from the horizontal, the image may be rotated using the Silent Vision Glasses Coupling Lens Assembly:

- On the coupling lens assembly, slightly loosen the rotation lock screw by turning it counterclockwise.

- The rotation lock screw may then be pushed to the left or right until the eye image on the monitor is correctly oriented.

- Re-tighten the rotation lock screw.

Real Eye Tracker Interface

Auditory Stimulus

Uses SoundBlaster or Presentation Software.

The auditory stimulus interface includes:

- Stereo amplifier.

- Binaural (stereo) auditory feed through a pair of low-distortion silicone tubes which have about a 10-11 msec delay and a bandwidth of 5kHertz.

- Foam ear inserts (in various sizes) to fit in the subject’s ears.

- Auditory tubing inserts.

- A sound level meter to calibrate sound delivered to the subject.

Foam Ear Tip with

with Auditory Tubing Insert &

Low-Distortion Silicone Tube

- Attach the short color coded Auditory Tubing Inserts onto the end of the Low Distortion Silicone Auditory Tubes which are hanging on each side of the Gantry mounting bar.  Place Red on the right side and the Blue on the left.

- Attach the Foam Ear Tip Insert into the Auditory Tubing Inserts. A connector (tube nibble) may be required to keep the insert from coming out.

- Slowly roll (rather than squeeze) the Foam Ear Tip into as small a

diameter as possible.

- Quickly insert the compressed plug well into the ear canal.

- Hold in the ear with finger tip until expansion is complete.

- Repeat process for the remaining ear.

- The tubing may be taped on the subject’s cheek to assure that

the insert does not become dislodged during testing.

- Click on the SoundBlaster or Presentation software from the desktop on the Stimulus Computer located at the operator’s console.

  - Adjust the sound intensity level (db).

  - Click on Start.

Insertion depth affects interaural attenuation.  Be sure to select the correct size Foam Ear Tip to obtain the best fit in the ear canal.  Insert as far as possible into the ear canal.

Somatosensory / EMG Stimulus

The Grass S88 Nerve / Muscle Stimulator features include:

- Two output channels for separate nerve locations (with independent control),

- Quartz crystal controlled digital timing circuitry,

- Variable pulse rates and durations,

- Variable train rates and durations,

- Variable delays, and

- Variable stimulus voltage.

Grass S88 Electrical Nerve/Muscle Stimulator

One Channel Single Repetitive Nerve / Muscle Stimulation

The Grass S88 consists of two stimulus channels:  S1 unit controls are silver; S2 unit controls are black.

| DIAL | Setting | MULTIPLIER SWITCH | Setting |
| --- | --- | --- | --- |
| Volt Control | Variable | Volt Multiplier | 10 SIU |
| Rate | Variable | Rate Multiplier | X1 |
| Duration | Variable | Duration Multiplier | X1 |
| Delay (ms) | *Optional | Delay Multiplier | *Optional |
|  |  | Function Switch | Single |

  - Turn on the Power for the S88. The switch is located at the bottom right corner of the unit. The orange light should turn on.

  - Turn ON the S1 UNIT.  The switch for the S1 (upper) unit is located in the right upper corner flanked by the ON pilot lamp and the red Overload Lamp.

- Press the ON Switch in the upward direction.

  - Set the S1 Volt Multiplier Switch to 10 SIU.

- Set the S1 Volt Control Dial to the lowest level. This will be equivalent to < 1 volt. (This is important because if the voltage is too high when the electrodes are plugged in, the subject might experience extreme discomfort).

  - Set the S1 Rate Multiplier Switch to X1.

- Select the S1 Rate (pulses per sec).

  - Select the S1 Duration Multiplier Switch scale.  (It is recommended that the multiplier is set on the X1 scale).

  - Select the S1 Duration (pulse length in ms).

  - Determine the S1 Delay Dial (ms) if any.

  - Set the S1 Function Switch to Single.

  - Reduce the skin impedance at the stimulus site using Nuprep Skin Prep.  Wipe the skin to remove the residual grit.

  - Apply Elifix Electrolyte to the electrodes.

  - Tape the electrodes securely to the skin.

  - Check the impedances using the Grass Impedance Meter. * Impedances should be < 5K ohms.  For further information refer to section entitled “Impedance Meter Check”.

  - Plug the Grass Stimulus Isolation Unit (SIU) into the cable labeled Grass S88 Stimulator – S1 (located on the front cabinet in the MSR).

  - Plug the stimulus electrodes (or electrode bar) into the SIU:

- Select the CURRENT RANGE (generally 1-15ma).

Stimulus Isolation Unit (SIU)

Select the POLARITY; generally normal.

  - Gradually increase the stimulus intensity using the Grass S88 S1 VOLTAGE CONTROL Dial.  Increase until motor threshold is reached (until the there is a visible muscle twitch) then back off to sensory threshold (the subject can feel the stimulus but there is no visible twitch).

*For subject comfort and maximal stimulation with the least amount of voltage move the electrode until it is positioned directly over the muscle/nerve while at motor threshold then back off to sensory threshold; make sure the skin impedance is  less than or equal to 5K ohms.

Response Pads

Lumitouch Response Pad

Lumitouch is a 5 key fiber optic response keypad. The system consists of absolutely metal free keypads connected via fiber optic cable to an opto-electronic controller unit. The cabling is fabricated from 100% plastic materials and therefore does not cause electromagnetic interference.

There are five response buttons which are arranged so that all or a subset of fingers can be used in a natural order and comfortable positioning when responding to a visual and/or auditory stimuli. Designed for placing on the arm rest.

- Instruct the subject in the use of the Lumitouch Response Pad.

FORP Response Pad

The FORP (Fiber-Optic Response Pad) is a fiber optic based response device.  The response pad is completely non-metallic and non-electronic.  It is intended for use in MEG, MRI or other environments where metal or electronics could hinder data acquisition, degrade image quality, or pose safety risks.

There are 4 colored buttons which are arranged in a diamond pattern.  It is designed to be held or cupped in the hand while using the thumb to respond to the appropriate color, position, number or pattern of the buttons.

- Instruct the subject in the use of the FORP Response Pad.

# NeuroScan Switch Response Pad

The NeuroScan is an electrical response device designed to use while lying on the arm rest to reduce the potential for arm movement.

There are four response buttons aligned for natural positioning of the fingers. There is an audible click with each button press.

- Instruct the subject in the use of the NeuroScan Response Pad.

MSR (Magnetically Shielded Room) & Subject Support

The Magnetically Shielded Room is designed to shield out or reduce the influence of environmental noise (i.e., the earth’s magnetic field, cars, elevators, etc) on the MEG signal.  The walls consist of a layer of a special metal which has a high magnetic permeability (Mu Metal) and two layers of copper.  Although the MSR greatly reduces the contribution of environmental noise, it does not adequately shield out lower frequency noise.

Room Lighting

Overhead Lighting

  - Turn the lights on in the MSR by depressing the square green button on the VAC Power Supply Unit which is located on the bottom shelf of the stand outside the MSR.

  - Dimming the lights in the MSR:

    - Use the Brightness (black) knob to dim the light.

Door Operation

Closing the Door:

- Push the door up to the door frame as far as possible.

- Press (until the air locks activate) then quickly release the OPEN/CLOSE BUTTON to seal the door.

Open/Close Button

Opening the Door:

- Press (until the air locks release) then quickly release the OPEN/CLOSE BUTTON (Black Button).

MSR Door

Pull the door open.

If the door fails to open, press the EMERGENCY OFF BUTTON (Red button) located on the outside and the inside of the MSR.

Emergency Door

Key

If the door fails to open when using the Emergency Off Button, then insert the EMERGENCY DOOR KEY in the upper and lower EMERGENCY UNLOCK (Yellow) keyholes and turn.  This will manually release the door locks.  Keys and keyholes are located on both sides of the MSR door.

# Positioning a Subject in the Chair

The gantry should be at 15o   tilt in the upright (sitting) position.

- Move the attached Visual Stimulus Screen up out of the way by loosening the black knob (located on the gantry mounting bar) before placing a subject in the chair.

- Plug the Head Coil Cables into the appropriate color-coded locations on the Head Localization Module.  *Route the right Head Coil Cable around the back of the subject’s neck so that it does not lie across the subject’s chest.  Route the other two cables in front of the subject’s neck so that they do not touch the subject’s chest.

| Head Coil Cables |  |  |
| --- | --- | --- |
| Nasion | Right Ear | Left Ear |
| Blue | Red | Green |

Head Localization Module

Note:  If the head coil cables are placed in the wrong jack or placed on the wrong anatomical location the coil locations will be skewed and the calibration will fail.  Repeat the calibration after changing the jacks the correct location in the jack box or switching the coils to the correct location on the subject.

- Mounting the Electrode Jack Box onto the amplifier when performing EEG:

- Touch the Static Discharge Button on the Amplifier Unit.

- If the Gain Calibration Card (Shorting Plug - small green circuit board) is situated on the Amplifier Unit, carefully remove it before connecting the Electrode Jack Box.

Gain Calibration    Card (Shorting Plug)

- Connect the Electrode Jack Box to the amplifier Unit.  Align the corner with the brackets on the amplifier.  Press down firmly.

Electrode Jack Box

Amplifier Unit

Never check impedances using the Grass Impedance Meter when the jack box is connected to the MEG system amplifier.

- Place the Stabilizing Head cuff:

- Open the valve stem and squeeze all the air out of the head cuff so that it is completely deflated.

- Close the valves.

- Place the protective head cover (cap) over the subject’s head.

Protective Head Cover

Place the Stabilizing Head cuff on the subject’s head so that the red cross mark is at the center front of the forehead and the connecting tubing is at center back.  *Make sure the bladders are positioned so that they lie flat inside the headband.

- Pull the head cuff over the ears.

- The “Y” shaped top piece should be placed so that the cross bar is at the back of the head and the other two arms of the “Y” are at the sides of the head.

- Adjust the top straps with the Velcro so that the headband does not cover the eyes.

- Place the hoses so that they run down the back of the neck; behind the chair.

- There are two pump valves; one to inflate the front of the bladder and the other to inflate the side bladders.

Head cuff Pumps

- Instruct the subject to hold the pumps which will be inflated when the subject is properly raised up and inserted into the dewar.

Stabilizing Head cuff

- Place a rolled up towel or pillow under the knees and so that the subject does not slide down in the chair.

- Make CHAIR ADJUSTMENTS:

The chair is a pneumatic-hydraulic unit and there may be a 2-3 second delay when turning controls from the unlocked to the locked position or when making chair adjustments.

There are two kinds of controls: The hand wheel which controls the directions and rate of movement and the valve which controls which adjustment is selected or activated.  Turn the hand wheel and valves slowly when making any adjustments.

After making  any adjustment using the hand wheel or valve allow the air pressure to equalize.

Hand Wheel Control

## Chair Movement / Positioning

The BRAKE valve controls the back and forward movement of the chair; located on the side panel.

- Slowly turn the brake valve to the unlocked position.

- Physically push the chair back or forward so that the subject’s head is positioned directly under the gantry; the back of the head should at the back of the dewar.  Check the subject’s posture.  Ask the subject to scoot their hips all the way back in the chair so that they are sitting upright in the chair.

Chair Brake Valve

After adjusting the chair set the valve to the locked position to prevent movement of the chair.    The valve should always be in this position unless adjustments are being made.

*Check to make sure the optical cables are not on the track before moving the chair.

## Chair Backrest Tilt Adjustment

The upper valve controls the TILT of the chair backrest; located on the front panel.

- Slowly turn the Tilt Control Valve to the unlocked position.

- Adjust the backrest so the subject is sitting upright in the chair and their head is positioned directly over the hips.  It is important make sure the subject is comfortable and therefore reduce the potential for fatigue.

- Turn the hand wheel:

- Clockwise  to make the backrest move forward.

Tilt Control Valve

Counter-clockwise  for backward motions.

- Turn the Tilt Control Valve to locked position.  It should always be in this position unless adjustments are being made.

## Chair Height Adjustment

The lower valve controls the HEIGHT of the chair; located on the front panel.

- Slowly turn the Height Control Valve to the unlocked position.

- Position the subject so that his/her head is at or near the top of the dewar. *When near the top of the dewar periodically ask the subject to raise himself/herself up in the dewar to determine how much room is remaining. Do not raise the chair while the subject is rising up in the chair.

- Turn the Hand Wheel:

Height Control Valve

Clockwise  to make the chair move upward.

- Counter-clockwise

for downward motions.

- Turn the Height Adjust Valve to

locked position unless adjustments

are being made.

* Care should be exercised when adjusting the chair height while a subject is seated under the gantry.  Until completely familiar with the operation of the chair adjust the chair very slowly in small increments.

- Instruct the subject to inflate the Stabilizing Head cuff to a comfortable level after positioning the subject’s head into the dewar helmet.

- The subject’s head should be centered in the helmet.

- Instruct the subject to squeeze both pumps at the same time to inflate the front and side bladders simultaneously until his/her head feels comfortable (moderately confined) in the helmet.

- When both bladders are inflated fold the tubing and clamp at fold so that air can not leak out.

- Move the hand pumps to a location near the wall as far away from the dewar helmet as the tubing allows. Place the pumps so that they are stationary and there is no movement in the clamp or the pumps.

- Tape the Head Coil Cables to the side of the dewar. *If applicable, tape the EEG Electrode Cables to the side of the dewar as far away from the Head Coil Cables as possible.

- Place a Chin Head Movement Monitor using tape to give the subject feedback regarding head movement.  Attach tape to each side of the dewar. The tape should just barely touch the subject’s chin.  *This step is optional.

- Adjust the stimulus.

- Instruct the subject regarding the task.

- Allow the subject to practice the task.

- Begin recording.

*When done recording deflate the bladder by opening the pump valves.  Some external pressure may have to be applied to deflate the air bladders fully.

Also move the chair forward, place the chair backrest in the upright position.

# Installing the Bed for Supine Recording

# The gantry should be at a 90o   tilt for supine recording.

*Never operate the gantry while a subject’s head is in the helmet.

- Raise the screen.

- Lower the Gantry to 90o tilt (horizontal position).  Gantry Controls are located on the right side of base. The controls require the operator to use both hands to operate.  If either handle is released the gantry motion will stop.

- Raise the Brake Control Lever (lower / long handle).

- Then raise the Direction Control Lever (upper / shorter handle).  This will release the brake.

- Then move the Direction Control Lever in the downward direction.

- Using the Angle Indicator readout located on the gantry near the front pivot point position the gantry at 900.

- Release the Direction Control Lever to stop the motion.  Release the Brake Control Lever.

- Slide the chair forward.

- Recline the chair backrest to approximately a 45o tilt.

- Remove the arm rest from the chair.  The back cushion may be removed or left in place.

- If the bed will have its angle adjusted place the Roller Wheel Assembly.  If the chair back cushion has been left in place, first slide the backrest cushion to its lowest position. Mount the Roller Wheel Assembly to the exposed seat back support plate.

- Using two people, lift the bed by the ends using the hand grips which are mounted to the under side of the bed.

- Set the bed onto the unpadded armrests of the chair. Align the holes.

- Insert the four mounting bolts, which are tied to the base of the bed, through the holes in the bed base into the holes in the chair armrests.

  - All four bolts should be screwed in a couple of turns and

  - fully tightened when all four are engaged.

- Install bed side rail on one side.  One side rail must be in place before allowing a subject to get onto the bed. Tighten rail clamp to secure.

- Raise the bed so that it is level with the dewar.

Bed Side Rail

Slide the bed back so that it rests against the dewar.

|  | *Lift the bed above the optical cables and any other sensitive equipment. |
| --- | --- |

# Positioning a Subject on the Bed

- Place a sheet over the bed. Place a towel at the head of the bed.

- Cover the Foam Pad Pillow and place inside the dewar.

- Place the step stool on the opposite side of the rail to assist the subject in getting onto the bed.

- Place a protective cap over the subject’s head.

- Have the subject sit on the edge of the bed then lie down.  Care should be taken so that the subject does not bump his head on the screen or the dewar.

- Ask the subject to bend his knees and use his legs to scoot himself up so that his head is inside the dewar.

- Place a towel or pillow case over the subject’s head to restrict head movement.

- Place a rolled up towel or pillow under the subject’s knees so that there is no pressure exerted on the back, tailbone, or knees.

- Place the head coil cables so that they do not lie across the subject’s chest. Plug into the jack box.

- Once the subject is on the bed, place the side rail on the remaining side.  Tighten the rail clamp to secure.

- If necessary, make sure the ancillary mirror has been placed in the 2nd (middle) wall mirror slot.

- The LCD projector image should be reversed / inverted to allow for the extra mirror bounce:

From the LCD Projector  touchpad (or using the remote):

- Press MENU.

- Press  /  to select the “PRJ Mode”.

- Press  /  to select desired projection mode (Rear / Reversed image).

- Press MENU to exit.

- Re-position the ceiling mirror.  Adjust the wall mirror.

- Begin Data Acquisition.

- When finished, remove the bed, move the chair forward, place the backrest in the upright position, replace the arm rest and back cushion.

- Position the gantry at 15 degrees.

# Subject Monitoring Inside the MSR

The operator can converse with and monitor the activities of the subject inside the MSR by:

- Voice Intercom / Parabolic Microphone / Flat Speaker

- Infrared Lighting for video in darkness

- Video Display at the operator console

- Video Cameras

- VCR for recording the video

Acquisition Workstation / Operator’s Console

- Video Display Monitor: Used to monitor a subject inside the MSR.  Located at the acquisition workstation / operator’s console.

  - Press the on/off switch is located on the front panel.

  - Press the channel button once to access the next channels.

  - Continue to press to access other channels.

| Channel 1 | Camera 1 (Infrared Camera) |
| --- | --- |
| Channel 2 | Camera 2 (Color Camera) |
| Channel 3 | VCR Play back (connected to Color Camera) |
| Channel 4 | (Blank) |

*Always monitor the subject via the Video Display Monitor when running a study.

THE SUBJECT MUST BE OBSERVED AT ALL TIMES DURING TESTING.

- Voice Intercom:  Used to communicate and listen to a subject inside the MSR.  The Intercom Electronics Box is located on the acquisition workstation / operator’s console. (There are no intercom controls inside the MSR, just a parabolic microphone and flat speaker).

Voice Intercom

  - Press the Red power button to turn on the Intercom Electronics Box to listen to a subject inside the MSR.

  - To speak to a subject who is inside the MSR press the PUSH TO TALK button located on the front of the Intercom Electronics Box. (Release the button to talk to the subject.)

  - Adjust the volume in the MSR or at the Console by using the volume control knobs located on the front of the Intercom Electronics Box.

- Infrared Lighting:  Illuminates a subject inside the MSR when recording in a dark environment so that the subject may be seen via the Video Display Monitor.    The Control box is located on the stand outside the MSR.

Infrared Lighting Control Box

Press the Red On/Off switch to turn on the power for the Infrared Light Power Supply (SEC 1212).

- Video Cameras: Used to view a subject who is inside the MSR on the Video Display Monitor. Located inside the MSR; reposition manually.

- Press the channel button located on the Video Display Monitor to change from Camera 1 to Camera 2.

| Channel 1 | Camera 1 (Infrared) |
| --- | --- |
| Channel 2 | Camera 2 (Color) |

Video Cameras & Microphone

*See Video Display Monitor for details.

- VCR:  The Sony VCR Located at the acquisition workstation / operator’s console. Use the VCR to record a video of a subject during an experiment. The VCR has the ability to send time codes to the rack along with the data so that the video can be referenced during data analysis.

VCR Panel Closed

VCR Panel Open

  - Make sure Acq has been launched,

  - Turn on the VCR by pushing the power button located on the top left of the panel, just below the word “Sony.”

  - Click on Windows > Collection Parameters.

  - In the Collection Parameters window under the General section select the Time channel check box.

  - Put an appropriate description in the Run Description window (i.e. with VCR Recording).

  - Setup the stimulus and instruct the subject.

  - When ready to start recording, click on the Acquire Data button at the bottom of the Run Description window and the Acquisition Monitor will come up.  The Time display box at the right side of the plot window will be blank.

  - Open the front panel by pulling the front of the VCR down and towards you.

  - Insert a tape in the appropriate slot.  The Time display box will become active and the time displayed will match the time display on the VCR.

  - Stop the tape and rewind to the beginning of the tape.

  - Do the head localization and prepare to start recording.

  - Press the record button  to begin recording and verify that the Time display box matches the VCR.

- The time information will be stored as a separate channel in the MEG/EEG dataset.

- The time data can be displayed in the plot window by selecting TIME in the Channel Selection dialog window.

- This should be a series of upward steps occurring once per second, visible if you set the vertical scale to 1000.  At minute and hour boundaries there will be a very large step in the data.

- The Sony video equipment does not output the video frame number as part of the time code, so resolution less than a second has to be calculated based on the number of samples since the last step in the TIME data channel.

  - To stop recording, press the stop button.

  - Rewind by pressing the   button.

  - Playing Back: The monitor has four channels that can be changed by pressing the Channel button on the video display monitor. The VCR channel is number three.

  - Eject the tape by pressing the         button.

  - When finished, close the panel.

# Recording Data / Data Acquisition

# Electronics / System Characteristics

The electronics system is responsible for processing MEG and EEG data including amplifying, filtering, sample rate conversion and firmware gradient formation.   It facilitates the measuring and recording of weak magnetic fields and electric fields produced by the nervous system simultaneously on many channels and provides sufficient processing power for user required algorithms to be implemented.

The electronics rack consists of the MEG DSQ bank and amplifiers and the EEG bank and amplifiers and the Peripheral Interface Unit (PIU) which includes the ADC Channels Interface. They are housed on the electronics rack.

System characteristics:

- 275 MEG Channels

- Total of 64 EEG Channels (56 Unipolar EEG Channels & 8 Bipolar EEG Channels)

- 16 ADC Channels

- 8 Bits; 16 Trigger Channels

- 28 Reference Channels

- Signal detection down to DC (0 Hz)

- Dynamic range: 32 bits

- Linearity better than 10-6

- Sampling rate: up to 4000 Hz. *higher sample rate are possible for subsets of all channels.

# MEG Electronics Rack

# ADC Setup:

The Analog to Digital Converter is a device that converts the analog signal to a digital signal that represents equivalent information.  The ADC Channel Interface is used to connect the input from various stimulus / stimulus response equipment to the electronics the signal can be recorded, displayed and stored along with the data.  The ADC channel interface is located on the front panel of the electronics rack.

The ADC interface connections should always be checked prior to any collection.  If using a stimulus response / reply system (Lumitouch, FORP or NeuroScan), before starting acquisition the user should check to make sure that the ADCs are connected to the proper response buttons.  Ask the subject to push each button individually.  Visually inspect the ADCs on the video display monitor to assure that the ADCs are connected to the correct response button in accordance with the research study design.

There are 16 ADCs.  The head localization coils are internally connected to ADCs 1-3.

ADC Channels Interface Boards

The Lumitouch Response Pads are often plugged into ADCs 9-15.

Lumitouch Reply System Interface Box

The FORP is usually plugged into ADCs 5-8.

FORP Interface Box

The auxiliary BNC cable can be used for other stimulus equipment or other computer configurations.

Auxiliary BNC Cable

# Data Acquisition

The main function of Acq is to set collection parameters such as sample rate, collection duration, set patient information, set data filters, set synthetic gradiometer order, set adaptive balancing, set up triggers and ADCs, collect and save data, perform head localization to co-register the sensors to the head coordinate frame, save collection parameters for future use.  Collection Parameters are set via dialog boxes prior to recording and stored in a protocol file.  However, these parameters can be adjusted for a given study without permanently changing the stored protocol file and should always be reviewed before starting a collection.

Data can be collected in one of three modes:  1) Continuous recording displays all data immediately while recording; 2) Triggered recording stores individual or; averaged sets of trials with specified time duration around some form of stimulus or external trigger; 3) Un-triggered recording stores trials of a specified duration in response to the operator’s command without stimulus triggering.

The final preparation before recording is the localization of the subject’s head with respect to the standard coordinate system.  This is done by localizing the position of coils placed on the fudicial landmark points of the subject’s head via the SQUID sensors after the subject’s head is placed in the dewar helmet.

From the Acquisition Console / Workstation:

- From the NIMH MEG Core Facility Acquisition Manager:

  - Select User

  - Launch Acq

NIMH MEG Core Facility Acquisition Manager

- Select Study.

- To view/edit Collection Parameters:

  - Click on Windows

  - Click on Collection Parameters

Collection Parameters Window

- Click OK.

- At the Acq Dialog Window enter:

Acq Dialog Window

  - Enter Subject ID (subject initials & MIS#).

  - Enter Procedure Description.

- Click on Acquire Data.

- Create Directory? Click Ok.

- When performing EEG, an IMPEDANCE CHECK should be performed on the MEG system:

- Select EEG Setup Window menu.

- Click on CHECK IMPEDANCE.  The results are displayed in the appropriate columns.

  - To check the impedance of all active EEG channels press the CTRL key on the keyboard at the same time as clicking on the impedance check button.

  - Please note: The impedance check requires 3 electrodes: The electrode being measured and at least 2 other electrodes.

- Click OK / EXIT.

- Click on HEAD LOCALIZATION button located on the side bar.

- Instruct subject for task and Head Localization. If applicable UNBLANK SCREEN.

- At the Stimulus Workstation: Click the mouse to activate the “LOCALIZING HEAD” prompt.   *If there is no head localization prompt, omit this step.

- Click on OK at the Head Localization Starter TO START THE LOCALIZATION.

Head Localization Window

- Click on CONTINUE to accept the head localization results. (acceptable tolerance is < 0.5 mm)

- STARTING THE STIMULUS:

  - If acquisition is NOT TRIGGERED click mouse on the Stimulus

Computer simultaneously with the start of the Data Collection (Step# 15),

- If trial is TRIGGERED click mouse after the start of collection

(Step#  15).

- To start DATA COLLECTION click on START.

Acquisition Monitor

- To select the ADCs and Channels to be displayed during collection: Click on the CHANNEL SET BUTTON which is below the group of Channel Scale Indicators.

Note: Select Custom… to customize the channel selection using the Channel Selector dialog.

Note:  Fluctuating Red lights in the Collection Status Window during acquisition   indicates that the data is being buffered; a failure condition is indicated only if the collection status buffer in completely red (full).

Solid yellow lights on the DSQ cards during acquisition indicate those channels are actively acquiring data; not a failure condition.

- Click on the CHANNEL SELECTOR DIALOG;

a.   From the UNSELECTED Window click on the

channels (i.e., Stim, ADC’s, EOG, EEG and/or  MEG channels) to be displayed.

b.   Click on the ARROW  to transfer those channels to the SELECTED window.

*Make sure the ADCs have been set up correctly and document.

c.   Click on APPLY.

Channel Selector Dialog

- Adjust channel (display) scale.

- When done acquiring (at the end of the trial) when data collection stops, head localization will start automatically. Wait for the HEAD LOCALIZATION PROMPT.

- To Print the Head Localization results: When Head Localization is done Click at the bottom of the screen and the Printer Icon will display. Click on the printer icon then click anywhere within the Head Localization dialog.

- Click on “Continue” to accept the data.  *If localization fails click on the REPEAT button to re-try localization.  At the General Information window, click OK.

- When head localization is done instruct the subject to relax and blank the screen.

- To begin another trial :

  - Click on FILE.

  - Click on CLOSE WINDOW.

  - Acquisition Manager window displays.

  - Edit the trial information in the Procedure area.

  - Click on OK.

  - Repeat  steps  5 - 22.

- When done inform subject of study completion and blank the screen.

- To end recording session click on CLOSE window.

- EXIT Acq.

- Select Dataset to save.

- Click on OK.

- From the Acquisition Manager window select the dataset to push to the storage RAID SERVER (tako.nimh.nih.gov).

- Click on Push Data.

NIMH MEG Core Facility Acquisition Manager

# Transferring Data To DVD

The DVD burner is located on the analysis computer -Tako.  Data must be pushed from the Acquisition Computer (Squid) to the Analysis Computer (Tako) and then copied to DVD.  If the data has already been pushed to the RAID server (as outlined above) skip to step 4.

- Log onto the Analysis Computer (Tako).

a.  	Login to Tako using your username and password.

b.  	Open a terminal window.

c.  	Type: “cd data”.

- Make sure that the staging area is empty.

- Type: “ls /tako_data1/dvd”.

- If the staging area is not empty

Type: “rm –r /tako_data1/dvd/(directory name)/”.

- Copy the data into the staging area.

- Type:  “cp –a (directory name)  /tako_data1/dvd”.

- Write data to DVD.

  - Insert DVD.

Type: “burndvd”.

- Unmount volume.

a.   	Right click on cdrom icon (located on the Start Here menu).

- Click on Unmount.

- Delete (remove) the data from the Analysis Computer (Tako) staging area  only:

a.  	Type: “rm –r /tako_data1/dvd/(directory name)/”.

(*this deletes the data from the staging area only; the data will  remain on the server for several weeks before being deleted.)

|  | When Someone Is Logged Into Tako Occasionally when trying to use tako to burn a DVD someone else may be logged in. In this situation, start a new xterm window by clicking on the terminal icon on the bottom of your screen (it looks like a picture of a computer monitor). Then make yourself the user of that window using the following command: > su - userName |
| --- | --- |

# Transferring Data To A Remote Computer

Tako, the analysis computer, is set up so that the data directory in your home directory is located on the RAID Array.  So it is appropriate for the data to get pushed there.

# After pushing data to the RAID array it can be acquired by logging onto Tako (tako.nimh.nih.gov) using your username and password.

  - Move to the data directory, i.e.,

“> cd data “

$Home/data directory structure:

date I subject_studyType_date_run.ds I your study data

  - Copy the data to the Remote Workstation for access.  This can be done in two ways.

If logged into Tako

    - From your $HOME/data directory type in:

“> scp -rp (date of study in year/mo/day format) username@yourWorkstation.ninds.nih.gov:.”

Example:

To copy data that was acquired on January 22, 2003 to a remote workstation called synapse.ninds.gov type the following command from your $HOME/data directory:

“scp -rp 20030122 userName@synapse.ninds.gov:.”

If logged into a Remote Workstation

a.	From your $HOME/data directory type in:

“> scp –rp username@tako.nimh.nih.gov:data/(date of study in year/mo/day format).”

Example:

To copy data that was acquired on January 22, 2003 from the MEG workstation called tako.nimh.nih.gov type the following command from your $HOME/data directory:

“> scp -rp userName@tako.nimh.nih.gov/20030122 .”

*Do not change the name of the MEG data during transfer or by using the “mv” command. If you need to change the name of a data file, use the following command: > changeDsName oldName.ds newName.ds

# Clean Up Procedures

MEG

Subject Clean Up

- If applicable, remove the EOG / EMG electrodes.

- Remove the black marker dots using an alcohol wipe.

- Clean the skin with a soft cloth and water; rub gently.

Disinfection Procedure

- Soak EOG / EMG electrodes in a freshly made 0.5% solution of Hypochlorite or approved hospital grade disinfectant for 10 minutes.

- Rinse thoroughly and allow to air dry.

Equipment Clean Up

- Clean THE HEAD LOCALIZATION COILS with alcohol.

- Clean the EOG / EEG electrodes with distilled water and a mild detergent.

a.	Gently rub the electrodes using a toothbrush to dislodge electrolyte residue.

  - Rinse well.

  - Pat dry using a soft towel.

  - Store in a dark dry environment.

|  | *When done recording slide the chair forward. Also lower the chair to the floor (so that the bellows is about 2 inches deep); adjust the chair backrest so that it is completely upright (at approx 150 angle). |
| --- | --- |

EEG

Subject Clean Up

- When recording is finished, remove the cap (or electrodes).

- If PASTE was used, take a damp cloth and gently rub the scalp in all electrode locations.

- If COLLODION was used, apply a solvent to dissolve the glue.  Use in a well ventilated area.

- Instruct subject to wash hair as usual with shampoo.

Disinfection Procedure

- If using the EEG CAP,

- Remove the electrodes from cap.

- Soak the ELECTRODES and EEG CAP in a 5% Sekusept-Plus solution or other hospital grade disinfectant for 10-15 minutes.  Afterwards clean electrodes as described below.

- If using a COLLODION / ELECTRODE application,

- Soak electrodes in a freshly made 0.5% solution of Hypochlorite or approved hospital grade disinfectant for 10-15 minutes.

- Rinse thoroughly and allow to air dry.

Equipment Clean Up

- Clean ELECTRODES with distilled water and a mild detergent.

- Gently rub electrodes using a toothbrush to dislodge electrolyte residue.

- Rinse well.

- Pat dry using a soft towel.

- Store in a dark dry environment.

- Clean the EEG CAP with warm tap water and a mild detergent.

- Rinse well.

- Pat dry using a soft towel; allow cap to air dry.

- Clean HEAD LOCALIZATION COILS with alcohol.

Appendix

In The Event Of An Electronics/Software Variance The Following Procedures Should Be Followed:

- Head Localization Failure:

  - Check Collection Parameters.

  - Check head movement.

  - Check head coil placement.

  - Repeat head localization.

- Acq Aborts (crashes):

  - Check front panel for electronics failure. Record codes.

  - If no electronics failure then re-launch Acq.

  - Check collection parameters.

  - Re-enter Acq information.

  - Fill out a Variance Report.

3.	Boot Electronics Failure:

a.   No LED’s – no power. If there is no fan noise or clicking of relays then it is likely      that there is a problem with the power.  Contact MEG Core staff for assistance.

b.   No LED’s – no 12 kHz Sync signal is being received by the card. Exit Acq then     turn the power off and on and try again. Without the 12 kHz sync pulse all the cards in the sub-banks will be held in reset.  Check the DSQ 2041 (the source of the sync signal).  Check the DSQ2047 bridge cards and fibers.  Check the DSQ 2043 is configured correctly for DSQ back or processing bank operation.    If the problem persists contact MEG Core staff for assistance.

- Flashing yellow – Exit Acq then turn the power off and on and try again. Indicates a failure in the self diagnosis.  If the problem persists contact MEG Core staff for assistance.

- Solid yellow DSQ2045 - Exit Acq then turn the power off and on and try again.  Indicates a problem in the self diagnosis of DSQ2045.  If the problem is persistent contact MEG Core staff for assistance.

- Solid yellow DSQ2010 and/or DSQ2043 - Exit Acq then turn the power off and on and try again. Indicates that there is no clock signal supplied to a bank (via the DSQ2041 and DSQ2047 fiber link).If the problem is persistent contact MEG Core staff for assistance.

- An entire bank or even several cards within a bank fail to boot - check the power supplied to the bank. Contact the MEG Core Facility staff for assistance.

- Slow boot – if a DSQ2043 takes 60 sec to boot rather than the regular 12sec, this is because the data fiber optic link between banks is broken or somehow disconnected.

- Flashing Green Lights – if green lights still flashing after launching Acq  then exit Acq. Turn the power off  for  5-10 minutes then turn on and try again.

- Report all occurrences; Fill out a Variance Report.  Follow-up with the MEG Core staff.

Contact MEG staff for assistance if unable to resolve problem(s).

Rebooting / Powering the Rack Off & On After a Failure

The status of the power is indicated by the On/Off Light Indicator (lamp) adjacent to the pushbutton.  When the green light is lit that indicates the power is ON. The electronics must be on and ready for use before running Acq.

- TO TURN THE POWER OFF - Depress the green ON/OFF Pushbutton.

- Wait about 5-10 minutes before depressing the green ON/OFF Pushbutton again to turn the system back on.

- AFTER TURNING THE RACK ON - wait for the system to reach thermal stability, that is, until all yellow lights on the front faceplate of the electronics cabinet are OFF and all green lights are ON before running Acq.

- Green LED Indicators - located on the front faceplate of the plug-in units in the electronics cabinet indicate the status of the individual units in the electronics.  For normal operation all of the green LEDs on the front panel should be on.

- Yellow LED Indicators - located on the front face plate of the plug-in units in the electronics cabinet indicate a fault has been detected in the unit; when lit there is a fault condition.  The system will not boot properly when a fault condition exists.

- Login as root.

Type:” # insmod qla 2X00”.

- Select user then launch Acq.

INDEX

A

Abralyt Skin Prep	15

Acq Aborts (Crashes)	52

Acq Dialog Window	46

Acquisition Console	43

Acquisition Monitor	50

Acquisition Workstation	43

ADCs	23,43, 44, 45

ADC Channel Interface	43

Adhesive Washers	6, 7, 9, 10, 12

Amplifiier	31

Analog to Digital Converter	43, 44

Analysis Computer	50, 51, 52

Ancillary Mirror	20, 38

Applying Electrolyte	15

Auditory Foam Ear Tip Inserts	25

Auditory Stimulus	25

Auditory Tube Nibble	25

Auditory Tubing Inserts	25

Avotec Real eye Tracking System	18, 24

Avotec Silent Vision	18, 24

B

Bed Side Rail	37, 38

Binaural Auditory Silicone Tubes	25

Binocular Fiber Optic Video Presentation System	24

Bipolar Channel	17

Black Screen Function	22

BNC Cable	44

Boot Electronics Failure	52

Brake Valve Control	33

C

Cap, Protective	32

Ceiling Mirror	20

Chair Adjustments	31, 33, 34

Chair Backrest Tilt Adjustment	34

Chair Height Adjustment	34

Chair Movement / Positioning	33

Channel Selector	45, 46

Channel Set Button	48

Chest Belt, Easy Cap	12

Chin Head Movement Monitor	35

Clamp, Headcuff	35

Clean Up Procedures	49, 50

Collection Parameters	43

Collection Status Window	48

Continuous Recording	45

Custom Button	50

Cz Placement	10, 11

D

Data Acquisition	45

Data Collection	48, 49

Degaussing	5

Delay, Stimulus	26

De-Metaling	5

Digital Camera	6, 8

Digitizing Electrode Positions	13

Disinfection Procedure	53, 54

Door Operation	29

Duration, Stimulus	26

Dynamic Range	43

DVD, Transfering Data To	50

E

Easy Cap System	10, 54

EEG Easy Cap System	10

EEG Electrode Impedance	7, 9, 10, 16, 45

EEG Bank	43

EEG Channels	43

EEG Setup	10

Electroculogram (EOG) Eye Lead Placement	7, 9

Electrode Impedance Meter Check	  16

Electrode Impedance Check , MEG System	47

Electrode Jack Box	16, 17, 31

Electronics Rack	43

Electrolyte	18

Elifix Electrolyte	7, 9, 10, 15, 16

Emergency Door Key	29

Emergency Off Button	29

Emergency Unlock	29

EOG	See Electroculogram

EOG Impedance	7, 9

Equipment Clean Up	49, 50

F

Fiber Optic Response Pad	28

Fiducial Points/Head Coil Placements	6, 8

Fit Test	6, 8

Foam Ear Insert Tips	25

Poam Pad Pillow	38

FORP Response Pad	28

FP1 Placement	10, 11, 12

FP2 Placement	10, 11, 12

Fpz Placement	10, 11

Function Switch, Stimulus	26

Flashing Green Lights	52

Flashing Yellow	52

G

Gantry	6, 8, 36, 38

Gantry Limits	6, 8

Gantry Mounting Bar	21, 25

Grass Impedance Meter	6, 7, 9, 10

Grass S88 Nerve / Muscle Stimulator	26

Grass SIU Unit	27

Grass Stimulator Duration Multiplier / Dial 	26

Grass Stimulator Multiplier Switch	26

Grass Stimulator Rate Multiplier / Dial	26

Grass Volt Control Multiplier / Dial 	26

Grass Stimulator Dely Multiplier / Dial	26

Ground Electrode	17

Gain Calibration Card	31

H

Hand Wheel Control, Chair	33, 34

Head Circumference	7, 8, 10, 12

Heal Localization Coils	7, 9, 35, 44, 53, 54

Head Localization	, 31, 45, 47

Head Localization Failure	52

Head Size/Fit Test	6, 8

Height Adjust Valve, Chair	34

Horizontal Recording	36

HP 715 Digital Camera	6,8

Head Stabilizing Cuff	32

I

Illumination Sensor, Optical	23

Impedance	5, 7, 9, 16, 17

Impedance Check, Meg System	44

Impedance Meter Check	16

Infra Orbital	7, 9

Infrared Eye Position Tracking System	24

Infrared Light Power Supply (SEC 1212) 	24

Infrared Light	29

Inion	6, 8, 10

Installing The Bed For Supine Recording	 36

Intensity, Stimulus	27

Intercom Electronics Box	40

J

Jack Box	16, 17

L

LapTop Computer	13

LCD Projector Remote Control	22

LCD Projector System, Sharp	19, 23, 38

Lens Adjustments	20

Lens Colllar	20

Light, Infrared	29

Light, MSR	29

Linearity	43

Linked Electrode Adaptor	17

Linked Mastoids	17

Localization	46

Lumitouch Response Pad	28, 44

M

Magnetically Shielded Room	See MSR

Mastoid Electrode Impedance	10

Mastoid Placement / Reference Electrode	10, 17

Measuring / Marking Frontal Polar & Occipital   Locations	10, 11

MEG Channels	43

MEG DSQ Bank	43

MEG Electrode Impedance Check	47

MEG Server	47

Mirror Adjustments	20

Motor Threshold	27

Mounting The Cap	12

MSR	5, 29

Multiplier Switch, Stimulator	26

N

Nasion - Inion Distance	6, 8, 10

Nasion	6, 8, 10

Neuroscan Switch Response Pad	28

No Led’s	52

Nuprep Skin Cleaner	7, 9, 10, 15

Nerve / Muscle Stimulator	26

O

One Channel Single Repetitive Stimulation	26

Orbital Ridge	9

Overhead Lighting	29

Oz, Marking & Measuring	10, 11, 12

O1, Maring & Measuring	10, 11, 12

O2, Marking & Measuring 	10, 11, 12

Operator’s Console	24, 25, 38, 39, 40

Optical Illuminaton Sensor	23

Outer Canthus	7, 9

P

Parabolic Microphone	39, 40

Peripheral Interface Unit	See PIU

PIU	43

Polhemus Digitizer	13

Polhemus Stylus Pen	13, 14

Polhemus Transmitter Module	13, 14

Polhemus Receiver	13, 14

Positioning A Subject In The Chair	31

Positioning A Subject On The Bed	38

Positioning The Projector	20

Powering the Rack	57

Preauricular Point	6, 8

Presentation Software Program	8, 25

Protective Cap	32

Projector, Sharp LCD	19, 20, 22, 37, 38

Pushing Dataset To The Meg Server	46

R

RAID Array	50, 51, 52

Rate, Stimulus	26

Real Eye Tracking System	 24

Rebooting	57

Receiver, Polhemus	13, 14

Recording Data / Data Acquisition	43

Reducing Skin Impedances	15

Reference Channels	43

Reference Electrodes / Mastoid Placement	10

Remote Control, LCD Projector	 22

Remote Workstation	52

Response Pads	28

Room Lighting	29

S

Sampling Rate	43

Screen Adjustment	14, 19, 20, 21, 22, 36, 37, 45, 46

Seat Recording	310

Sensory Threshold	27

Sharp LCD Projector System	18, 19, 38

Shorting Plug	31

Side Rail	36, 37

Signal Detection	43

Silent Vision Video Presentation System	24

Silicone Tubes, Auditory	25

Single Repetitive Stimulation	26

SIU	See Stimulus Isolation Unit

Skin Prep	7, 9, 10, 15

Slow Boot	52

Solid Yellow Dsq2010 And/Or Dsq2043	52

Solid Yellow Dsq2045	52

Somatosensory / EMG Stimulus	26

Sony VCR	41

Soundblaster	25

Sound Level Meter	25

Squid	51, 52

Stabilizing Head Cuff	32

Step Stool	38

Stereo Amplifier	25

Stimulus Computer	47

Stimulus / Response Equipment	28, 44

Stimulus Isolation Unit	27

Stimulus Screen	21

Stimulus Software	25

Stimulus Workstation	45

Stylus Pen, Polhemus	13, 14

Subject Clean Up	49, 50

Subject Monitoring Inside The MSR	39, 40, 41

Subject Prep	5

Supine Recording	36

T

Tako	50, 51

Telescoping Screen Arm	21

Tilt Adjust Valve, Chair	34

Time Display Box	41

To Begin Another Trial	47

Touch Pad, Response	28

Touch Pad, LCD Projector	22

Tragus	6, 8

Transfering Data To A Remote Computer	51

Transmitter, Polhemus	13, 14

Trigger Channels	43

Triggered Recording	45

Tube Nibble, Auditory	25

U

Unipolar Channel	17

Upright Seated Recording	33, 34, 35

Untriggered Recording	45

V

Valve, Chair	33, 34

Variance Reporting	38, 39

VCR, Sony	41

Video Display Monitor	39

Video Recording	41

Visual Interface	18

Visual Software	18

Visual Stimulus Screen	21

Visual Stimulus	18

Voice Intercom	40

Voltage Control, Stimulus	26

W

Wall Mirror	20

Y

Y Connector…………………………………………..17

Z

Z Position	17